ABSTRACT

Ghifari, Fauzan. Utilization of  Barium Sulfate (BaSO4) Contrast Media as

a  Replacement  for  Radiograph  Marker.  Scientific  Writing,

Radiodiagnostic and Radiotherapy D3 Study Program. STIKes Widya Cipta

Husada Malang. Advisors: (1) Yeni Cahyati, S.Si., M.Si. (2) Deni Dwiyanti,

S.Si

Contrast  media  is  a  material  that  can  be  used  in  radiological  examinations

which aim to provide differences in the density of the surrounding organs. One

of  the  contrast  media  used  in  radiological  examinations  is  positive  contrast

media,  namely  barium  sulfate  (BaSO4). Marker  is  a  device  made  of  material

with a high atomic content used as a marker or information on the radiograph

by placing it outside the object being examined. The purpose of this study was

to  determine  whether  the  remaining  positive  barium  sulfate  (BaSO4)  contrast

media  could  be  used  as  a  marker  substitute.  The  research  design  used  was

experimental  research  (true  experiment).  This  research  was  conducted  by

drying  barium  sulfate  (BaSO4),  then  making  a  marker  and  then  printing  a

marker.  The  research  procedure  used  3  variations  of  the  exposure  factor,

namely the low, medium, and high exposure factors and density testing (black)

using  a  densitometer.  The  conclusion  obtained  is  that  the  remaining  contrast

media can be used as a substitute for markers with an average density value of

0.423 for low voltage, 1.123 for medium voltage, 1.303 for high voltage. This is

in accordance with the reference value for optical density of 0.2-2.5.

Keywords: Positive contrast media, Barium sulfate (BaSO4), marker